All mirrors will be used at 1031nm under 1.146° incident angles. All coating must be optimized for this wavelength and incident angle.

All mirrors must have chamfer of 0.5mm orientated by 45° on the front side and on the back side.

- A set of few plan/plan ULE substrates (quantity will depend on unit price):

  - Diameter ~1 inch, thickness ~6.35 mm. A measurement with uncertainties of the thickness and diameter is mandatory

  - Front side:

    - Super polished <0.1nm RMS (measured)

    - CA=80%

    - Scratch dig ≤ 10-5

    - Flatness ≤ 0.10 wave P-V @633nm

    - HR coating transmission T≤2ppm: T must be as closed as 2ppm with the constraint that the coating must be an exact quarter wave stack. Measurement must be provided.

    - Coating absorption < 1ppm (nominal value 0.5ppm) and coating scattering <5ppm (as small as possible). Measurements must be provided.

  - Back-side :

    - Scratch dig ≤ 20-10

    - Flatness ≤ 0.25 wave P-V @633nm

  - anti-reflection coating

- a set of few plan/concave ULE substrates (quantity will depend on unit price):

  - Diameter ~1 inch, thickness ~6.35 mm. A measurement with uncertainties of the thickness and diameter is mandatory

  - Radius Of Curvature (ROC) of the concave surface after HR and AR coatings: 2241.0 mm (and 500 mm for half of the set). For practical reason we would need to know the uncertainty on the expected ROC.

  - Front side:

    - Super polished <0.1nm RMS (measured)

    - CA=80%

    - Scratch dig ≤ 10-5

    - Flatness ≤ 0.10 wave P-V @633nm

    - HR coating transmission T≤2ppm: T must be as closed as 2ppm with the constraint that the coating must be an exact quarter wave stack. Measurements must be provided.

    - Coating absorption < 1ppm (nominal value 0.5ppm) and coating scattering <5ppm (as small as possible). Measurements must be provided.

  - Back-side :

    - Scratch dig ≤ 20-10

    - Flatness ≤ 0.25 wave P-V @633nm

  - anti-reflection coating

3) A set of few plan/plan sapphire substrates (quantity will depend on unit price):

- Diameter ~1 inch, thickness ~3.0 mm. A measurement with uncertainties of the thickness and diameter is mandatory

- Front side:

    - Super polished <0.1nm RMS (measured)

    - CA=80%

    - Scratch dig ≤ 10-5

    - Flatness ≤ 0.10 wave P-V @633nm

    - HR coating transmission T~120 ppm: T must be as closed as 120 ppm with the constraint that the coating must be an exact quarter wave stack. Measurements must be provided.

    - Coating absorption < 1ppm (nominal value 0.5ppm) and coating scattering <5ppm (as small as possible). Measurements must be provided.

- Back-side :

    - Scratch dig ≤ 10-5

    - Flatness ≤ 0.25 wave P-V @633nm

    - Back-side parallel to front-side within 60 arc seconds

- anti-reflection coating